RITHI

|  | Function | Name | Date | Signature |
| --- | --- | --- | --- | --- |
| Written by |  |  |  |  |
| Checked by |  |  |  |  |
| Checked by |  |  |  |  |
| Approved by |  |  |  |  |

CHANGE HISTORY for template DVP_FO_51

NB: This history is in “masked text” and will not be visible when the document is printed. It can be deleted when this template is used for a project.

| Version | Date | Change | Author |
| --- | --- | --- | --- |
| 1 | 05/03/2004 | Creation and distribution for comment | B. MARCHAL |
| 2 | 01/01/2011 | Merger of specifications Proposed use included Addition of an appendix for the expression of requirements | T. BOULANGER |

CHANGE HISTORY

| Revision | Date | Change | Author |
| --- | --- | --- | --- |
| 00 | 01.06.2022 | Initial | Devika |
| 01 | 01.03.2023 | Revision for “Daily Call Review” Validation | Devika |
| 02 |  | Re-validation for RITHI v2 (Supabase): all processes, the security scan of 27-Oct-2026, and the sections re-filled from the RITHI validation package |  |

TABLE OF CONTENTS

1  Description of the Software

1.1  Intended Use

1.2  Scope

1.3  References

1.4  Definitions and acronyms

2  Process Definition

2.1  Process needs to be followed as per International Standards

2.2  Software Tool(s) usage in the Process

2.2.1  Assumptions and dependencies of software tool

2.2.2  Risks associated with each activity

2.2.3  Storage of data

3  Validation Planning

3.1  Type 1 - In-house Developed software

3.2  Type 2- OFF The Shelf (OTS) Software

3.3  Type 3 – Disposal

4  Risk Analysis and Control Measures

5  Validation Tasks

6  CONCLUSION

# Description of the Software

RITHI software is an internally developed solution for the service process of customer-related activities — installing, maintaining and complaint handling of all our ventilators in the field through their life cycle. RITHI v2 (the version this revision validates) replaces the AppSheet application of Revision 01: a web application whose records are held in a managed PostgreSQL database (Supabase), with every access rule enforced in the database itself.

RITHI CRM records and manages field-service operations for medical devices: service/complaint-adjacent call registration, engineer visit reports, spare-part requests and approvals, stock movements, preventive-maintenance scheduling, warranty/contract cover and customer feedback. It is used within the QMS to create and retain quality records; it is NOT medical-device software and does not perform a medical-device function.

Software category: Predominantly GAMP Category 5 (custom application) running on Category 1/4 infrastructure & configured platform services (Supabase/Postgres, GitHub Pages).

## Intended Use

Service operation and administration as follows:

•	Sale entry and ownership transfer

•	Complaints handling, daily complaint review and feedback

•	Installation and preventive maintenance

•	Contract management

•	Spare management and engineer hand stock

•	Indoor (workshop) service

•	Training records and controlled documents

•	Service KPIs and quality objectives

## Scope

All ALMS supplied Products (Imported and Indigenous). The processes in scope, each validated below:

•	1. Order processing – Spare request

•	2. Stock maintenance for service spare parts

•	3. Sales – Warranty and Non-warranty

•	4. Annual Maintenance Contracts

•	5. Complaint handling

•	6. Installation & Preventive maintenance

•	7. Customer feedback

•	8. Indoor service (workshop)

•	9. Masters, users and access

•	10. Training and controlled documents

•	11. Reports, KPI and quality objectives

•	12. Data storage, audit and security

In all, 72 screens, each assigned to one of these processes; the requirement behind each screen is in the RITHI validation package [4].

## References

Reference documents:

[1] IEC/TR 80002-2:2017 — Validation of software for medical device quality systems.

[2] Service Procedure QSP-05; Customer feedback QSP-08.

[3] Medical Devices Rules, 2017 (CDSCO) — Fifth Schedule (QMS); ISO 13485:2016 §4.1.6 (validation of QMS software); ISO/TR 80002-2:2017; GAMP 5 (2nd edition); Computer Software Assurance (CSA) — risk-based validation effort; ISO 14971:2019 (risk management); Information Technology Act, 2000 (legal standing of electronic records).

[4] RITHI validation package, VAL-RITHI-CRM version 3.2 (DRAFT) — held in RITHI under Software Validation: user and functional requirements, the traceability matrix (docs/REQUIREMENTS.md), risk register, test protocols and defect register.

[5] Validation run record (docs/VALIDATION_RUN.md) — the automated execution of every database test suite and check; last run 27-Oct-2026.

[6] Security Scan report (docs/SECURITY_SCAN.md) — penetration test and vulnerability scan, 27-Oct-2026.

[7] Servicing requirements against ISO 13485 §7.5.4 (docs/ISO13485_SERVICING.md); Cover requirements (docs/COVER_REQUIREMENTS.md); Call Request requirements (docs/CALL_REQUEST_REQUIREMENTS.md).

[8] RITHI handbook (docs/HOW_TO_USE.md).

## Definitions and acronyms

| Initials and acronyms | Meaning |
| --- | --- |
| MD | Medical Device |
| RM | Risk Management |
| PRM | Preliminary Risk Management |
| URS / FRS | User Requirement Specification / Functional Requirement Specification |
| IQ / OQ / PQ | Installation / Operational / Performance Qualification (test protocols) |
| NAR | Non-Auditable Requirement — asked for by the system owner, not offered as evidence against a regulatory clause |
| RLS | Row-Level Security — the database decides which records each signed-in person may read or change |
| RBAC | Role-based access control — Roles & Permissions |
| UCN | Unique Call Number |
| DCCR | Daily Complaint Review Register (R/SER/35) |
| FFR | Field Failure Report |
| SA / MC / OT | Sale entry / Maintenance Contract / Ownership Transfer numbers |
| WGP / OGP / AMC / CMC | Within Guarantee Period / Out of Guarantee Period / Annual Maintenance Contract / Comprehensive Maintenance Contract |
| PM | Preventive Maintenance |
| MRN / MTN | Material Return Note / Material Transfer Note |
| GAMP | Good Automated Manufacturing Practice (ISPE) |
| ALCOA+ | Attributable, Legible, Contemporaneous, Original, Accurate — and complete, consistent, enduring, available |

# Process Definition

All process as per Service Procedure QSP-05, Customer feedback QSP-08, such as:

•	Complaints registration, daily review and solution reporting

•	Preventive maintenance activities

•	Customer feedback collection

•	Installation activities

•	Spare requests, issue and engineer hand stock

Software tool RITHI covers all above process mentioned in QSP-05 and QSP-08, and the further processes listed in 1.2.

## Process needs to be followed as per International Standards

Refer QSP-05 and QSP-08. The servicing obligations of ISO 13485:2016 §7.5.4, and the clauses it reaches into, are assessed requirement by requirement in [7]; the validation of the software itself answers §4.1.6 and IEC/TR 80002-2 [1].

## Software Tool(s) usage in the Process

1. Order processing – Spare request

Ability to place a request for spares from engineers, against a call or for hand stock; Line-by-line approval by the Reporting Manager, Commercial and NSM as the cover requires; Re-assignment of a request before it is issued, with a log. Screens: Spare Requests, RM Approval.

2. Stock maintenance for service spare parts

Stock out by stores to the corresponding engineers, with the Delivery Challan; Acknowledged receipt per delivery; Hand stock derived from issues, consumption, transfers and returns; Consumption booked against a visit; reconciliation and voiding. Screens: Pending Dispatch, Stock Out, Hand Stock, Hand Stock Report, Stock Transfer (MTN), Material Returns (MRN), Spare Consumption, Part Master, Part Search.

3. Sales – Warranty and Non-warranty

Product sale entry with its warranty; The Product Database names the party of the latest sale entry or ownership transfer; Ownership transfer of a machine. Screens: Warranty Register, Product Database, Product Database 2.0, Product Master (product lines), Ownership Transfer, Product & Party Search, Machine History.

4. Annual Maintenance Contracts

Contract entry, renewal and PM visit count. Screens: Contract Register.

5. Complaint handling

Breakdown call registration (from a call request, or directly); Daily complaint review stage 1; Call allocation; Daily complaint review stage 2; Visit work details entry; Service report entry with spares used; Service manual report upload; Daily complaint review stage 3; Field Failure Report raised from a review. Screens: Request Registration, Pending Registrations, Field Call Register, Pending Calls, Daily Complaint Review Register (R/SER/35), Call Review, Visit Reports / Service Reports, Bulk Report Mapping, Field Failure Register, Product Failure Analysis, Solved Without a Report.

6. Installation & Preventive maintenance

Installation call generation; PM call generation (bulk and PM Due); Installation & PM call visit details and closure report. Screens: Installation Calls, Preventive (PM), PM Bulk Upload, PM Due, Machines Without an Installation Call.

7. Customer feedback

Collection of feedback. Screens: Customer Feedback, Feedback Without a Report, Reports — Customer Feedback Report.

8. Indoor service (workshop)

Intake, cleaning, work, quality check and release of equipment in the workshop; Pre-Delivery Testing and recycled spares. Screens: Indoor Service Register, Spare Recycling, Pre-Delivery Quality Check.

9. Masters, users and access

Party, user and master-list maintenance; Roles & permissions; a login only for a person on the User Master; Signatures, held for and placed by their owner. Screens: Party Master, User Master, New Employee Requests, All Masters, Roles & Permissions, Admin Config, SLA / Objective Configuration, Settings, Device Cache Status.

10. Training and controlled documents

Training assignment, sessions and Training Form R/HR/11; QMS documents and service manuals. Screens: Training, QMS Documents, Service Manuals, Technical / Service Notes, How RITHI Functions.

11. Reports, KPI and quality objectives

KPI, objectives and dashboards computed by the database; Reports and exports; Bulk uploads of historical data. Screens: Dashboard, My Workload, KPI & Failure Analysis, Objective, Spare Insights, Tracker, Bulk Uploads, Data Export, Reports, Reports — Consumption Report, Reports — KPI Export, Reports — Not Consumed Against this Call, Reports — Call Report, Reports — Stores Dispatch Report.

12. Data storage, audit and security

Audit log and record retention; quality records never deleted; Row-level security; nobody signed out or unknown holds anything (security scan); Backups and validation records. Screens: Audit Log, Software Validation, Version History.

### Assumptions and dependencies of software tool

Assumptions:

•	User should have computer knowledge and training in RITHI.

•	Each user signs in with their own login, made only for a person on an active User Master row; the role on that row decides what they may do.

•	Users have internet access; engineers in the field may work from the copy of the machine register kept on their device, which refreshes itself when the server changes.

Dependencies:

•	Supabase — Managed PostgreSQL, Authentication, PostgREST API, Edge Functions, backups (criticality High).

•	GitHub (Actions + Pages) — Source control, CI build (tsc + Vite), static hosting of the web client (criticality Medium).

•	Resend (optional) — Transactional email for the daily digest and the scheduled table export (neither deployed as at this revision) (criticality Low).

### Risks associated with each activity

| Process ID | Process | Risk Identified |
| --- | --- | --- |
| 1 | Order processing – Spare request | R-67: A request is raised against the wrong machine, or for a part that does not exist in the master R-68: A line passes a stage on an answer that does not mean yes, or skips a review it needed R-69: A spare is rejected or dropped, or an order moved to another engineer, with no reason R-70: A batch approves a line past a review, or a line the approver may not decide R-71: A spare is cleared for Stores with no grounds, or under a contract number that is not one R-72: The name shown against an approval, a drop, a receipt or an office consumption line is the one the client sent … and 5 more (section 4) |
| 2 | Stock maintenance for service spare parts | R-74: The document for a shipment cannot be produced again, or lists other than what the shipment carried R-76: Anybody holding the transfer permission moves anybody's stock, to any name R-77: A movement is back-dated across a closed period, or dated in the future R-79: A rename moves records belonging to somebody else, leaves the person's own behind, or rewrites a signature R-84: A part is renamed in the catalogue only, merged into another, or mapped so it is not offered on the calls it fits R-87: A transfer or recovered entry is recorded against the wrong machine among those sharing a serial, or cannot be saved at all … and 15 more (section 4) |
| 3 | Sales – Warranty and Non-warranty | R-54: A register fails or is cut short and the history is shown as complete R-85: The assembled record is read stale as current, rebuilt by somebody unauthorised, or read by people not entitled to the warranty and contract detail it assembles R-86: A Sale Entry, Contract Entry, machine, transfer or additional entry is deleted R-88: A QMS entry's number, revision, effective date or file is overwritten in place, or a controlled document is a link that can change R-107: A loaded or reported installation moves a machine's warranty to the wrong dates, or another machine's R-41: A register has no key its interface can infer, so a second load adds the file again … and 4 more (section 4) |
| 4 | Annual Maintenance Contracts | — |
| 5 | Complaint handling | R-44: The auto review answers a call it must not R-45: A YES raises no report, two reports, or a report carrying a CAPA decision nobody made R-46: A load attributes old judgements to the person loading them, raises reports for them, or cannot be told apart from reviews made here R-47: A bulk action includes a first-year or already-answered call, or an automatic save commits a value nobody chose R-48: The repeat-failure verdict, the age at failure or the visit context is wrong or silently absent R-49: A failure is counted against the wrong product, or the correction alters the call … and 19 more (section 4) |
| 6 | Installation & Preventive maintenance | R-56: A request is reported saved with some of its calls unwritten, or dated to an attendance later than today R-65: PM calls are created with no serial, with uncontrolled values, or twice from one file R-06: A call is filed under the wrong type R-08: Queries time out as volume grows |
| 7 | Customer feedback | R-62: A visit is dated in the future or before the complaint, or is saved without the parts fitted on it R-63: A role refused the export right downloads data as a workbook R-64: Feedback on calls a reader may not see is readable through the data interface |
| 8 | Indoor service (workshop) | R-66: A recovered part, an accessory or a check verdict is deleted R-29: Equipment held in the workshop cannot be identified, or its condition on arrival is unknown |
| 9 | Masters, users and access | R-80: A starting or reset password stays known to somebody other than its owner R-81: A person holds a permission or sight of every record that their role does not show R-82: A person's R&R history is deleted with their directory row, or a Department or manager name is recorded wrongly with nothing said R-83: A customer is renamed, or a verification re-signed, through the data interface R-97: A save rewrites roles nobody touched, narrows the administrator, or leaves no record of what changed R-100: An SLA target, the Hotline desk or a frequent-failure rule is changed with no record … and 6 more (section 4) |
| 10 | Training and controlled documents | R-89: An attendance, assessment, score or R&R period is overwritten with no trace, or an assignment cancelled without authority or reason R-17: An engineer works to a superseded manual, or an uncontrolled QMS document is presented as current |
| 11 | Reports, KPI and quality objectives | R-53: A workbook is downloaded by a role refused export R-55: A monthly figure is typed, recalculated or deleted with no record of who or what it was R-90: A file leaves in a format the export gate does not check, or its audit entry records a file that was never produced R-91: A load writes records the operator did not approve, or leaves no trace of who ran it R-96: Whole tables leave without the export authority, unrecorded, or cut short without saying so R-26: A view is rebuilt without `security_invoker`, so it reads as its owner and the policies beneath it stop applying |
| 12 | Data storage, audit and security | R-98: A person granted the right to read the trail sees nothing, or the trail is left in a shared browser after sign-out R-101: A re-test overwrites the earlier result, keeps the first date, or names somebody who did not execute it R-103: Records remain on a device after sign-out, or a preview writes under the administrator’s identity R-43: A fragment of the credential is written to a log the organisation does not control R-10: Malicious HTML stored in a KB article R-12: An untested change reaches production … and 4 more (section 4) |

### Storage of data

All data is stored in the cloud: the RITHI database (Supabase PostgreSQL), with uploaded files — service reports, delivery challans, attachments — in Google Drive.

•	Complaints history, visit reports and the daily complaint review

•	Customer feedback

•	Products sold, warranty, contract and ownership records

•	Spare requests, dispatches, consumption and hand stock

•	Installation and maintenance records

•	Training records, QMS documents and the audit log

Quality records are held in the managed Supabase PostgreSQL project, which provides automated backups (and point-in-time recovery on the applicable plan).

Recovery objectives are defined by the customer: RPO (max acceptable data loss) = ______ ; RTO (max acceptable downtime) = ______ .

Restore capability is verified periodically by restoring to a test project and confirming record integrity — a backup is not evidence of recoverability until a restore has been demonstrated.

# Validation Planning

|  | Module | Check plan | Responsibility |
| --- | --- | --- | --- |
| 1 | Order processing – Spare request | Ability to place a request for spares from engineers, against a call or for hand stock Line-by-line approval by the Reporting Manager, Commercial and NSM as the cover requires Re-assignment of a request before it is issued, with a log | Service engineer · Reporting Manager · Commercial · NSM |
| 2 | Stock maintenance for service spare parts | Stock out by stores to the corresponding engineers, with the Delivery Challan Acknowledged receipt per delivery Hand stock derived from issues, consumption, transfers and returns Consumption booked against a visit; reconciliation and voiding | Service Stores · Spare Coordinator · Service engineer |
| 3 | Sales – Warranty and Non-warranty | Product sale entry with its warranty The Product Database names the party of the latest sale entry or ownership transfer Ownership transfer of a machine | Service admin · Commercial |
| 4 | Annual Maintenance Contracts | Contract entry, renewal and PM visit count | Service admin · Commercial |
| 5 | Complaint handling | Breakdown call registration (from a call request, or directly) Daily complaint review stage 1 Call allocation Daily complaint review stage 2 Visit work details entry Service report entry with spares used Service manual report upload Daily complaint review stage 3 Field Failure Report raised from a review | Hotline · Customer Support Manager · Service engineer |
| 6 | Installation & Preventive maintenance | Installation call generation PM call generation (bulk and PM Due) Installation & PM call visit details and closure report | Service admin · Hotline · Service engineer |
| 7 | Customer feedback | Collection of feedback | Service engineer |
| 8 | Indoor service (workshop) | Intake, cleaning, work, quality check and release of equipment in the workshop Pre-Delivery Testing and recycled spares | Service engineer (workshop) · QC |
| 9 | Masters, users and access | Party, user and master-list maintenance Roles & permissions; a login only for a person on the User Master Signatures, held for and placed by their owner | Administrator · Service admin |
| 10 | Training and controlled documents | Training assignment, sessions and Training Form R/HR/11 QMS documents and service manuals | VP Technical · Engineer R&D and QMS |
| 11 | Reports, KPI and quality objectives | KPI, objectives and dashboards computed by the database Reports and exports Bulk uploads of historical data | NSM · Customer Support Manager |
| 12 | Data storage, audit and security | Audit log and record retention; quality records never deleted Row-level security; nobody signed out or unknown holds anything (security scan) Backups and validation records | Administrator · QA |

## Type 1 - In-house Developed software

YES. RITHI is custom software (GAMP Category 5), developed and changed under the project's own change control: every change is built, tested by the automated suites and recorded in Version History before it goes live.

## Type 2- OFF The Shelf (OTS) Software

The platform services RITHI runs on are configured off-the-shelf services (GAMP Category 1/4), assessed as suppliers in the validation package:

•	Supabase: Acceptable as the platform of record, provided the customer-side controls (RLS configuration, backup verification, access management) are maintained and evidenced. Leverage supplier certifications; qualify the configured use via IQ/OQ.

•	GitHub (Actions + Pages): Acceptable for source control and hosting the client. Recommend enabling branch protection on main and recording release approvals as part of change control.

•	Resend (optional): Low criticality; acceptable for operational notifications. Not in the GxP record path.

## Type 3 – Disposal

Quality records are never deleted by the application: a wrong entry is voided or cancelled with a reason and kept. Data is exported from RITHI to Google Drive by the scheduled exports (Data Export), and the audit log is retained and archived. On retirement of RITHI, a complete export of every register is taken and verified before the database is closed.

The following tables record the validation of the software before implementing, and the periodic validation every year and as and when required.

# Risk Analysis and Control Measures

Refer 2.2.2. Each risk of the RITHI risk register [4], with the control that mitigates it. "Validated By" is signed by the person verifying the control.

| Process ID | Process | Risk Identified | Mitigation | Validated By | Status |
| --- | --- | --- | --- | --- | --- |
| 1.1 | Order processing – Spare request | R-67 Raising a spare request: A request is raised against the wrong machine, or for a part that does not exist in the master | The call's identity and cover are copied from the call, not typed; a request raised from a call is fixed to it; the part is a pick list with no free text; the OR number is the database's |  | Residual Low · OQ-138 |
| 1.2 |  | R-68 Stage of a spare line: A line passes a stage on an answer that does not mean yes, or skips a review it needed | Whole-word answer rule in the database and the client, held together by check:ui; the stages needed are written once at RM approval; an auto-approval names nobody; spare_approval_whole_word_test and handstock_needs_nsm_test |  | Residual Low · OQ-139 — passed 27-Oct-2026 |
| 1.3 |  | R-69 Ending a spare line unsupplied: A spare is rejected or dropped, or an order moved to another engineer, with no reason | Required in the bulk function; TO BE required on the single-line, Pending Dispatch and reassignment paths (D-043) |  | Residual Medium · OQ-140 |
| 1.4 |  | R-70 Deciding many spare lines: A batch approves a line past a review, or a line the approver may not decide | One function decides each line at its own stage, one stage at a time; ineligible lines are skipped and counted; RM never per order; confirmation names the count; spare_bulk_approval_test and spare_bulk_decisions_test |  | Residual Low · OQ-141 — passed 27-Oct-2026 |
| 1.5 |  | R-71 Commercial and NSM clearance: A spare is cleared for Stores with no grounds, or under a contract number that is not one | Form rules per status and reason; only the Cleared status approves; the answer is retained in approval_data under the stage's permission; spare_approval_forms_test |  | Residual Medium · OQ-142 |
| 1.6 |  | R-72 Attribution of spare and consumption decisions: The name shown against an approval, a drop, a receipt or an office consumption line is the one the client sent | The dispatcher is stamped from the session (0211); the login is on every row in sys_updated_by (0244); the named columns TO BE stamped likewise (D-041, D-042) |  | Residual Medium · OQ-143 |
| 1.7 |  | R-73 Booking spares out of Stores: A stock out mixes engineers, is written in part, or sends more than is owed | One function, one statement, all or nothing; one engineer refused in the screen and the database; the quantity capped at the remainder; spare_dispatch_test, dispatched_by_stamped_test |  | Residual Low · OQ-145 — passed 27-Oct-2026 |
| 1.8 |  | R-75 Material returns: A return is recorded for another engineer by somebody not entitled, in part, or with no trail | Insert policy limits on-behalf returns to administrators and approvers; the stock guard; returns cannot be edited; TO BE written in one transaction and audited (D-044, D-051) |  | Residual Medium · OQ-153 |
| 1.9 |  | R-78 Saving a header and its lines: A failure between the header and the lines leaves part of the record | TO BE one transaction per record (D-044); the stock out is already one |  | Residual Medium · OQ-156 |
| 1.10 |  | R-105 Global search: The header search shows a record the person could not open on its register, or opens a different record from the one chosen | Every register is read as the signed-in person, so row-level security bounds the results exactly as on the register; a register is searched only where its screen may be opened; a hit opens by the record’s own key (UCN, UID, REQID, id); check:ui holds that every hit names a real page key and every screen reads the state its hit carries. OQ-212 |  | Residual Low · OQ-212 |
| 1.11 |  | R-04 Spare approval: A stage is approved by an unauthorised role | Per-stage DB guard; manager-scoped approval; OQ FRS-010/011 |  | Residual Low |
| 2.1 | Stock maintenance for service spare parts | R-74 Delivery Challan and Declaration: The document for a shipment cannot be produced again, or lists other than what the shipment carried | Lines read from what THIS stock out carried; signature only where the printer booked it; address changes limited to four columns and to dispatch; TO BE found at any age (D-045) |  | Residual Medium · OQ-146, OQ-147 |
| 2.2 |  | R-76 Stock hand-over: Anybody holding the transfer permission moves anybody's stock, to any name | The stock guard caps the quantity; TO BE limited to one's own stock unless authorised, and to directory engineers (D-049) |  | Residual Medium · OQ-154 |
| 2.3 |  | R-77 Dating of stock movements: A movement is back-dated across a closed period, or dated in the future | The stock-out movement is dated by the database clock; TO BE refused before the cutoff and in the future (D-050) |  | Residual Medium · OQ-155 |
| 2.4 |  | R-79 Correction of a person's name: A rename moves records belonging to somebody else, leaves the person's own behind, or rewrites a signature | Administrator only (database guard); one statement; blank, case-only and twin-name limits; each column matched as its own read policy matches; signature columns excluded by name; the refusing guards admit only a ticket no user can write; the move is stated before saving; the quality tables are imaged by record_audit; OQ-159, OQ-160 |  | Residual Low · OQ-159, OQ-160 — passed 27-Oct-2026 |
| 2.5 |  | R-84 Part identity and fit: A part is renamed in the catalogue only, merged into another, or mapped so it is not offered on the calls it fits | Rename only through rename_part(), one transaction across nine tables, merge refused, the consumption guard admitting only its ticket, moves stated first; bulk mapping written in small batches with every failure named; unrecognised products flagged; OQ-170, OQ-171, OQ-172 |  | Residual Low · OQ-170, OQ-171, OQ-172 |
| 2.6 |  | R-87 Ownership transfer and recovered cover: A transfer or recovered entry is recorded against the wrong machine among those sharing a serial, or cannot be saved at all | Model and serial required; owner moved only for that pair; previous owner left empty unless unambiguous; recorder stamped from the session; entries keyed on machine_key (OPEN until D-054 and D-060 are fixed); OQ-181 |  | Residual Low · OQ-181 |
| 2.7 |  | R-92 Loading historical servicing records: A loaded call, visit or spare record carries an attribution, a status or a parent the file does not establish | Registrant columns left empty where the file names nobody (FRS-051); visit key derived from UCN and date; status effect stated before confirmation; stubs marked Imported and counted; one visit per UCN from the file’s own date, no date meaning no visit; imported provenance on every migrated row. OQ-191 |  | Residual Low · OQ-191 |
| 2.8 |  | R-94 Loading historical stock movements: A movement file is loaded twice, or the import exemption reaches an entry that is not an import | The exemption keyed on source = import and on nothing else; every other check still applies; lines with no parent held back; a warning that a re-run adds rows until each register has a natural key. Residual Medium while D-071 is open. OQ-192 |  | Residual Medium · OQ-192 |
| 2.9 |  | R-95 Loading a master: A load renames or merges an entry, or leaves a value in place that the corrected file removed | Each master keyed on what identifies it; the Standard Complaint matched on its Key and never renamed; blanksClear only where the file is the whole row, and never on a stamped or required column; Party key assigned once. OQ-193 |  | Residual Low · OQ-193 |
| 2.10 |  | R-106 One-time data correction applied by a migration: A correction applied to the live record does part of its work, or changes a record’s identity without moving what names it | Identity changes go through the same rename a person uses, in one transaction, with every record moved; idempotent so a re-run finishes the work; the migration’s notices are printed in the run; a read-only probe is run before and after on the live project. D-077, D-079. OQ-213 |  | Residual Low · OQ-213 — passed 27-Oct-2026 |
| 2.11 |  | R-40 Completion of a value the record did not supply: The system fills a field with a value that is available but untrue, or discards the record for want of it | The completion is conditional on the source being able to answer; where it cannot, the field is left empty and the record kept; the invariant that made the fabricated value visible is retained rather than relaxed |  | Residual Low |
| 2.12 |  | R-34 Recording of parts fitted: A part fitted during a visit is not recorded against the call | The report writes the line still in the entry control, under the same validation as an added line; “Add Consumption?” must be answered Yes or None Consumed, and Yes requires a line; all lines are written in one statement |  | Residual Low |
| 2.13 |  | R-36 Reproduction of a personal signature: A signature is obtained or applied by somebody other than its owner | The signature is readable and writable only by its owner — no administrator path exists to read or set one; a document reproduces it only where the block names the person producing the document, on an exact identity match; removal by an administrator is a definer function that discloses nothing |  | Residual Low |
| 2.14 |  | R-07 Stock accuracy: Overdraw of stock through a hidden movement | Guard counts all movements irrespective of visibility; OQ FRS-013 |  | Residual Low |
| 2.15 |  | R-13 Hand-stock integrity: More of a spare is recorded as consumed than the engineer holds | Accepted by the user (2026-10-08) in exchange for never losing a fitted part: the line is booked and every active Spare Coordinator is notified at once with the engineer, part and quantities (0401); the visit form warns before saving; negative balances show on Hand Stock; transfers and returns still cannot exceed what is held; OQ-15 |  | Residual Medium |
| 2.16 |  | R-14 Reconciliation of consumption: An office user books, amends or voids a consumption entry without authority or without a reason, or alters which call/engineer it belongs to | consumption.reconcile permission enforced by RLS; mandatory UCN/engineer/part/reason in a DB trigger; identity fields immutable on amendment; original quantity, reason and author retained on the row; audit_log entry; OQ-16, OQ-17. RAISED FROM LOW ON 2026-09-05 when the independent before/after image (record_audit) was withdrawn; RETURNED TO LOW ON 2026-09-21 (0225), the image being taken again by a trigger that cannot be bypassed |  | Residual Low |
| 2.17 |  | R-19 Spare order reassignment: The engineer on a spare order is changed after the parts have been issued | Refused by the function AND by a table trigger, so no path is exempt; three independent tests for issue; every permitted change logged with both names, actor, time and reason; OQ-23 |  | Residual Low |
| 2.18 |  | R-21 Migrated vs. created data: A figure derived partly from migrated records is read as though the system produced all of it | The migrated arms are held separately and never merged; the balance reports the split and both readings, with the identity on_hand - hist_net = on_hand_live holding per row; OQ-30 |  | Residual Low |
| 2.19 |  | R-22 Period close: Closing a period changes the balance it was meant to preserve | The close and the arms divide the same instant from opposite sides, so no movement falls on both; verified neutral on production-scale data before release; OQ-31 |  | Residual Low |
| 2.20 |  | R-32 Condemnation and recovered parts: Equipment is scrapped without authority, or a recovered part re-enters stock as though it were new | Condemning is a permission of its own, enforced by a trigger and granted to the administrator alone on apply; the reason is required by a constraint. Recovered parts are recorded with a condition grade and credit NO stock balance, so no part can reach an engineer through this route at all. OQ-46 |  | Residual Low · OQ-46 |
| 2.21 |  | R-15 Partial issue & receipt: An outstanding balance is lost, or stock is credited that never arrived | Per-stock-out quantities held separately; the line stays queued for its remainder; receipt acknowledged per delivery and completed only in full; OQ-18 |  | Residual Low |
| 3.1 | Sales – Warranty and Non-warranty | R-54 Machine History completeness: A register fails or is cut short and the history is shown as complete | A failed read is named on screen and no count is called exact while one has failed; a capped read carries “+” and names the register. OPEN until D-019 is fixed |  | Residual Low · OQ-107 |
| 3.2 |  | R-85 Product Database 2.0: The assembled record is read stale as current, rebuilt by somebody unauthorised, or read by people not entitled to the warranty and contract detail it assembles | Live-as-of stamp and a five-minute rebuild when a source changed; rebuild refused without masters.edit, cover.edit or administrator; readership stated as a decision for the system owner rather than inherited from 0221; OQ-176 |  | Residual Medium · OQ-176 — passed 27-Oct-2026 |
| 3.3 |  | R-86 Retention of cover records: A Sale Entry, Contract Entry, machine, transfer or additional entry is deleted | No delete path and the 0049 retention guard on all six tables; no delete buttons; corrections imaged by record_audit (OPEN until D-055 is fixed); OQ-180 |  | Residual Low · OQ-180 |
| 3.4 |  | R-88 Control of QMS document revisions: A QMS entry's number, revision, effective date or file is overwritten in place, or a controlled document is a link that can change | Number, revision, effective date and file required and fixed once saved; a new revision is a new entry and the old one retired; number + revision unique; links flagged (OPEN until D-061 is fixed); OQ-182 |  | Residual Low · OQ-182 |
| 3.5 |  | R-107 Warranty start decided by the installation record: A loaded or reported installation moves a machine's warranty to the wrong dates, or another machine's | One row per call keyed on the UCN; the choice is the feedback's own two words; the start is the solved day in India time and the end the database's own period arithmetic; the result is written by the system and shown on the row; only the product + serial the row names is re-synced; loading needs Bulk Uploads; every one-time apply kept the old values |  | Residual Low · OQ-237, OQ-238 — passed 27-Oct-2026 |
| 3.6 |  | R-41 Re-loading a corrected export: A register has no key its interface can infer, so a second load adds the file again | Every loadable register carries a generated stored key with a plain unique index; the target is verified against a database by an automated check that refuses expression and partial indexes; a row without the key is refused and named |  | Residual Low |
| 3.7 |  | R-39 Keying of a device or document record: A record is keyed on less than what identifies it | The key is the model together with the serial, or the document together with the machine, held as a generated stored column; the sufficiency of each key was measured against the data before loading it |  | Residual Low |
| 3.8 |  | R-33 Device identification on a call: A call is raised without the machine’s serial number | The serial is a required field on the request and on every call form; the request refuses an item without one and names the master as the cause where it is absent; the request identifier carries the serial, so an unidentified call is visible in the key |  | Residual Low |
| 3.9 |  | R-11 Availability: Data loss without recoverable backup | Managed backups; periodic restore test; PQ/IQ evidence |  | Residual Medium |
| 3.10 |  | R-30 Decontamination: Equipment is opened or dismantled before it has been decontaminated | A database trigger refuses to record any recovered part while the job’s decontamination flag is false; the cleaning record names the work instruction AND its revision. OQ-44 |  | Residual Low · OQ-44 |
| 5.1 | Complaint handling | R-44 Automatic answer to Review 2: The auto review answers a call it must not — a first-year failure, an unknown age, a Review 2 a person had begun — or answers while nobody authorised has switched it on, or names the wrong person | The switch is held by review.auto alone and refused by the database to anyone else; the exclusions are in the database function, not the screen; each answer names the person who switched it on and carries an auto-review marker Review 3 reads; the run reports what it left for a person. OPEN until D-028 is fixed |  | Residual Low · OQ-84, OQ-85 |
| 5.2 |  | R-45 Field Failure Report raised by a review: A YES raises no report, two reports, or a report carrying a CAPA decision nobody made | A database trigger on the review raises it, from the base tables, at most once per UCN; it is never withdrawn; CAPA fields are left blank; imported reviews raise nothing. OPEN until D-029 is fixed for the CAPA defaults |  | Residual Low · OQ-86, OQ-87 |
| 5.3 |  | R-46 Loading historical reviews: A load attributes old judgements to the person loading them, raises reports for them, or cannot be told apart from reviews made here | The file’s names and dates are kept; no identity of the loader is stamped; every row is marked imported; no report is raised; matched on UCN so a re-load corrects. OPEN until D-029 is fixed |  | Residual Low · OQ-88 — passed 27-Oct-2026 |
| 5.4 |  | R-47 Bulk and automatically saved review answers: A bulk action includes a first-year or already-answered call, or an automatic save commits a value nobody chose | The database refuses first-year, unknown-age and answered calls in the bulk function; a confirmation states the count; pick lists never commit on a keystroke; every save is attributed to its session |  | Residual Low · OQ-89, OQ-90 |
| 5.5 |  | R-48 Evidence presented to the reviewer: The repeat-failure verdict, the age at failure or the visit context is wrong or silently absent | The verdict is computed in the database under both rules and states which fired; an unknown age reads not known, never a number; a failed context read says so |  | Residual Low · OQ-91, OQ-92 |
| 5.6 |  | R-49 Attribution of a failure to a product line: A failure is counted against the wrong product, or the correction alters the call | The actual product is held on the review, not the call; every analysis reads the review’s product where one is recorded; only review.edit holders may set it |  | Residual Low · OQ-93 — passed 27-Oct-2026 |
| 5.7 |  | R-50 Review completion dates: A completion date is changed by somebody other than an administrator, or without a record | The database refuses a non-administrator change and records every change with before and after, person and time. OPEN until D-020 is fixed |  | Residual Low · OQ-94 |
| 5.8 |  | R-51 Field Failure Report creation and completion: A report is raised without its problem or customer, renumbered, re-attributed, deleted, or left unreviewed with no means of recording the review | The number is issued and fixed by the database; the raiser is stamped and kept; mandatory fields refused by the database as well as the form; deletion blocked (0166); the weekly review recordable from the form. OPEN until D-027 is fixed |  | Residual Low · OQ-97, OQ-98 |
| 5.9 |  | R-52 Widened visibility for failure-report readers: The function that gives register readers the visits and spares behind a report returns other calls, or the printable report opens for somebody refused the register | The definer function tests ffr.view and returns only for a UCN that has a report; the print route checks the register’s module key before row-level security. OPEN until D-026 is fixed |  | Residual Low · OQ-99 — passed 27-Oct-2026 |
| 5.10 |  | R-57 Correcting a pending request: A correction writes values the request form would have refused, or changes who raised it | The correction is refused after Pending by screen and database (FRS-087); a zero-row save is reported as not saved. OPEN: every corrected field is free text and the address is correctable (D-030); OQ-112 |  | Residual Medium · OQ-112 |
| 5.11 |  | R-58 Disposition of a pending request: A request is mapped to a call that does not exist, reported closed when nothing was written, or left pending beside the call registered from it | Three stated outcomes; the row audit on call_requests records every change. OPEN: map-anyway confirm, no row count, back-fill swallowed (D-031); OQ-113 |  | Residual Medium · OQ-113 |
| 5.12 |  | R-59 Registering a call: A refused or failed registration is kept in the browser under a UCN of its own making | The database assigns the UCN (FRS-005) and refuses by RLS. OPEN: the local fallback (D-032); OQ-116 |  | Residual Medium · OQ-116 |
| 5.13 |  | R-60 Vigilance answers at registration: The three vigilance questions are recorded as NO without being answered | The section is emphasised; amendment needs its own right and is audited. OPEN: each field defaults to NO (D-033); OQ-117 |  | Residual Medium · OQ-117 |
| 5.14 |  | R-61 Amending, cancelling and re-opening a call: A sectional right cannot be used, a closed call is cancelled round the screen, or a re-open leaves no reason | Section rights enforced by triggers (0127); cancel needs calls.cancel and a reason; re-open needs pending.register or calls.create; record_audit captures the row. OPEN: D-035, D-036 (D-034 fixed 2026-09-30); OQ-120, OQ-123 |  | Residual Medium · OQ-120, OQ-121, OQ-122, OQ-123 |
| 5.15 |  | R-93 Loading historical Daily Complaint Reviews: The uploader is recorded as the reviewer, or a Field Failure Report is raised from a loaded review | The file’s reviewer names and dates kept; no session identity stamped on a loaded review; an import mark on every loaded review; the FFR trigger passes a marked review; keyed on UCN so a corrected file updates. Residual High until D-029 is closed — today a DCCR load would do both. OQ-88 |  | Residual High · OQ-88 — passed 27-Oct-2026 |
| 5.16 |  | R-99 Fleet frequent-failure rule: A fault repeating across machines of one model is not flagged, or repeated calls on one machine are read as a fleet fault | Counted on DISTINCT serials of one product within a window in days, the call under review counted; either rule marks a frequent failure and names which fired; recorded answers not re-based. OQ-198 |  | Residual Low · OQ-198 — passed 27-Oct-2026 |
| 5.17 |  | R-42 Configuration of a read-only role: A role described as read-only holds an authority that writes | The copy is a one-time seed rather than a standing mirror, so a later grant cannot cross; the write authority held by each such role is reported by diagnostic and presented as a matter for review, since a grant may be deliberate |  | Residual Medium |
| 5.18 |  | R-37 Attribution of a judgement on a failure: A completed review records no person, or records the wrong one | The identity is taken from the authenticated session at completion by a database trigger, on every write path including automatic and bulk saves; it is not re-stamped by later edits; the name is carried onto the report; no screen name is ever written |  | Residual Low |
| 5.19 |  | R-38 Amendment of a field failure report: A report is changed and the change is not recorded, or the record of it is altered | A database trigger records every UPDATE with its before/after values and its author, independently of the application; the log has no insert, update or delete policy and the deletion guard is armed on it; reading needs ffr.manage or administrator |  | Residual Low |
| 5.20 |  | R-03 Call status: Status does not reflect the latest visit | Trigger recomputes from latest entry; open_state generated column; OQ FRS-007 |  | Residual Low |
| 5.21 |  | R-05 Identifier assignment: Duplicate or missing UCN / Call Number | DB triggers + unique constraint + shared sequence; OQ FRS-005 |  | Residual Low |
| 5.22 |  | R-09 Bulk import: PM batch imports wrong/duplicate rows | Preview before commit; DB assigns identifiers; admin-only; OQ FRS-009 |  | Residual Low |
| 5.23 |  | R-18 Bulk re-allotment: A bulk edit writes rows the user did not intend, or allots outside the manager’s team | Selection is limited to the rows currently listed, never rows a filter hides; the edit touches one field; the engineer list is the reporting sub-tree and the write is constrained again by RLS; OQ-22 |  | Residual Low |
| 5.24 |  | R-20 Acting for a team member: A record is raised for an engineer outside the manager’s team, or its authorship is lost | The list offered is the reporting sub-tree; created_by retains the author independently of the engineer named; OQ-24 |  | Residual Low |
| 5.25 |  | R-27 Decision support: A suggested value is written as though a person had chosen it, or a value outside the controlled list is offered | The field is written only by the person choosing; the model re-ranks candidates and cannot return one outside them, enforced in the function AND in the client; what was offered and accepted is retained; OQ-36 |  | Residual Low · OQ-36 |
| 6.1 | Installation & Preventive maintenance | R-56 Raising a call registration request: A request is reported saved with some of its calls unwritten, or dated to an attendance later than today | One minted REQID and one insert for every call; Call Attended? must be answered and an Attended Date given on Yes. OPEN: the fallback path half-saves and no future-date refusal exists (D-030); OQ-111 |  | Residual Medium · OQ-111 |
| 6.2 |  | R-65 PM batch load: PM calls are created with no serial, with uncontrolled values, or twice from one file | Administrator only; preview; database assigns identifiers (R-09). OPEN: D-038; OQ-135 |  | Residual Medium · OQ-135 |
| 6.3 |  | R-06 Call-type routing: A call is filed under the wrong type | Routing triggers + per-table CHECK constraint; OQ FRS-006 |  | Residual Low |
| 6.4 |  | R-08 Data at scale: Queries time out as volume grows | Indexed once-per-query RLS; table split; PQ performance test |  | Residual Low |
| 7.1 | Customer feedback | R-62 Recording a visit: A visit is dated in the future or before the complaint, or is saved without the parts fitted on it | The date is refused by the form and by a database trigger for every form-entered visit (0115); a partial save names what is unsaved and Save retries only that; unaccounted parts are reported (FRS-055); OQ-126, OQ-127 |  | Residual Low · OQ-126, OQ-127 |
| 7.2 |  | R-63 Export gate: A role refused the export right downloads data as a workbook | csvExport checks export.data. OPEN: xlsxDownload does not (D-018); OQ-129 |  | Residual Medium · OQ-129 |
| 7.3 |  | R-64 Visibility of customer feedback: Feedback on calls a reader may not see is readable through the data interface | The module key; a browser filter. OPEN: fb_read carries no visibility rule (D-037); OQ-134 |  | Residual Medium · OQ-134 |
| 8.1 | Indoor service (workshop) | R-66 Workshop records: A recovered part, an accessory or a check verdict is deleted; cleaning is attributed to whoever the browser names; a change leaves no trace | Job rows cannot be deleted (no delete policy); QC and condemnation are rights enforced by triggers. OPEN: D-039; OQ-136 |  | Residual Medium · OQ-136 |
| 8.2 |  | R-29 Custody of equipment on the premises: Equipment held in the workshop cannot be identified, or its condition on arrival is unknown | A register with one row per intake carrying who received it, when, and its condition on arrival; `kind` distinguishes customer property from the organisation’s own stock on a separate axis from the work being done, so the distinction survives a change of activity; damage has its own field with the date it was reported to the owner. OQ-43 |  | Residual Low · OQ-43 |
| 9.1 | Masters, users and access | R-80 Issue and reset of login credentials: A starting or reset password stays known to somebody other than its owner | Creation and reset are administrator-only in the database; a reset is generated, shown once, ends every session and is recorded by the database without the password; the starting password is a common default its owner may replace and is not forced to — ACCEPTED by the system owner for ease of operation (D-057, 2026-10-01); OQ-161, OQ-162 |  | Residual High · OQ-161, OQ-162 |
| 9.2 |  | R-81 Access granted to one person: A person holds a permission or sight of every record that their role does not show | Individual grants writable by an administrator only; shown in the Access drawer; a clone states when it grants sight of every record; each grant, clone and role correction written to the audit trail; directory and sign-in role differences counted on screen; OQ-163 |  | Residual Medium · OQ-163 |
| 9.3 |  | R-82 The User Master record of a person: A person's R&R history is deleted with their directory row, or a Department or manager name is recorded wrongly with nothing said | Deletion refused where any profile, R&R or training record exists, and no cascade; every field the form offers saved; a manager name matching nobody stated before saving (OPEN until D-053 and D-059 are fixed); OQ-164, OQ-165 |  | Residual Low · OQ-164, OQ-165 |
| 9.4 |  | R-83 Customer identity, KYC attribution and the Serviceman: A customer is renamed, or a verification re-signed, through the data interface; or a Serviceman spelling matches nobody | The name refused on UPDATE by the database and never offered on screen; verifier and time stamped and discarded when supplied; Serviceman correction offers only User Master names, counts first and writes all or none (OPEN until D-058 is fixed); OQ-166, OQ-167, OQ-168 |  | Residual Low · OQ-166, OQ-167, OQ-168 |
| 9.5 |  | R-97 Changing a role’s permissions: A save rewrites roles nobody touched, narrows the administrator, or leaves no record of what changed | Only touched roles written; administrator always full; an empty set refused; the matrix read-only on admin.view; each change recorded with role, permission, old and new state, actor and time. Residual Medium while D-067 is open. OQ-196 |  | Residual Medium · OQ-196 |
| 9.6 |  | R-100 Changing a rule that decides quality answers: An SLA target, the Hotline desk or a frequent-failure rule is changed with no record | Database refuses a change from anyone but an administrator or config.manage; each change recorded with the old and new value, actor and time. Residual Medium while D-067 is open. OQ-199 |  | Residual Medium · OQ-199 |
| 9.7 |  | R-102 Admission and identity: A login with no profile, or a local demo account, is admitted with authority | Production identity provider only in a released build; an unresolved identity holds no permission in the application or the database; profile read failure signs out; one super-admin list agreed by client and database; every sign-in and refusal recorded. Residual Medium while D-074 is open. OQ-203 |  | Residual Medium · OQ-203 |
| 9.8 |  | R-104 The database a browser writes to: A browser is pointed at a database other than the released one | Change permitted to users.manage.settings only (a part of Manage users); a banner on every screen while connected elsewhere; each change recorded. Residual Medium while D-072 and D-067 are open. OQ-207 |  | Residual Medium · OQ-207 |
| 9.9 |  | R-35 Configuration of access roles: A role is created with no permissions and taken into use | A role can be created only by derivation from an existing role, so an empty set cannot be produced from the application; the key is slugged and cannot collide; roles are not deleted while users may be on them |  | Residual Low |
| 9.10 |  | R-01 Access control (RBAC/RLS): A user sees or edits records outside their authority | Server-side RLS keyed on auth.uid(); least-privilege roles; OQ tests FRS-003/004 |  | Residual Low |
| 9.11 |  | R-02 Audit trail: An action is not recorded, or a user alters the trail | DB-stamped identity; admin-only read; no user update; a bulk write is recorded as one attributable event rather than dropped; OQ test FRS-021 |  | Residual Low |
| 9.12 |  | R-28 Attributable registration: A call cannot be traced to the individual who registered it, or is traced to the wrong one | Two columns, not one: the desk the call is filed to, and the individual who registered it. The individual is taken from the authenticated session by a trigger and a caller-supplied value is DISCARDED, so it cannot be forged through the API; a supplied desk is accepted only when it names a hotline desk. Records with neither are shown as unattributed rather than attributed to anyone, and _registered_by_check.sql reports where that line falls. OQ-37 |  | Residual Low · OQ-37 |
| 10.1 | Training and controlled documents | R-89 Amendment of training records: An attendance, assessment, score or R&R period is overwritten with no trace, or an assignment cancelled without authority or reason | Cancellation restricted to training.manage in the database with a reason kept; amendments to sessions and R&R periods recorded with before value, author and time in a history nobody can edit (OPEN until D-062 is fixed); OQ-183 |  | Residual Low · OQ-183 |
| 10.2 |  | R-17 Document control: An engineer works to a superseded manual, or an uncontrolled QMS document is presented as current | Manuals are catalogued against the product and offered on the call itself rather than sought out; QMS documents carry number, revision and effective date; retiring withdraws a document from use while retaining it; the two shelves are maintained under separate rights enforced in RLS; OQ-21 |  | Residual Low |
| 11.1 | Reports, KPI and quality objectives | R-53 Downloads from analysis screens: A workbook is downloaded by a role refused export | The workbook writer tests the same permission as the CSV writer and the control is not offered; check:ui refuses a download function without the test; every download is audited. OPEN until D-018 is fixed |  | Residual Low · OQ-100 |
| 11.2 |  | R-55 Quality objective figures: A monthly figure is typed, recalculated or deleted with no record of who or what it was | Writing restricted to config.manage in the database; every change recorded by the database with before and after; deletion refused while figures exist; cut-offs lockable by an administrator. OPEN until D-021 is fixed |  | Residual Low · OQ-109 — passed 27-Oct-2026 |
| 11.3 |  | R-90 Taking a file out of the system: A file leaves in a format the export gate does not check, or its audit entry records a file that was never produced | One gate for every format (FRS-193.1); a report-level right in addition (FRS-193.2); an audit entry per file written only when the file is produced (FRS-193.4/.5). Residual stands at Medium while D-065 is open; Low once it is closed. OQ-186 |  | Residual Medium · OQ-186 |
| 11.4 |  | R-91 Bulk loading a register: A load writes records the operator did not approve, or leaves no trace of who ran it | Admin-only; preview of every row ready and held back; confirmation stating the count and the re-run behaviour; the preparation step writes nothing before confirmation; an audit entry per load; record_audit on the ten quality tables records the rows themselves. Residual Medium while D-067 and D-075 are open. OQ-189 |  | Residual Medium · OQ-189, OQ-190 |
| 11.5 |  | R-96 Whole-table export: Whole tables leave without the export authority, unrecorded, or cut short without saying so | Admin module; the tables list excludes the audit trails; read as the person with RLS in force; the export gate on the ZIP; an ordered paged read; a stated cut-off; an audit entry per ZIP. Residual Medium while D-065 is open. OQ-195 |  | Residual Medium · OQ-195 |
| 11.6 |  | R-26 Views and row-level security: A view is rebuilt without `security_invoker`, so it reads as its owner and the policies beneath it stop applying | check:views fails on any view over an RLS-protected table that lacks the setting; _status.sql reports it on a live project; OQ-35 |  | Residual Low · OQ-35 |
| 12.1 | Data storage, audit and security | R-98 Reading the audit trail: A person granted the right to read the trail sees nothing, or the trail is left in a shared browser after sign-out | The screen gate and the database policy name the same audience; cached rows removed at sign-out; record_audit readable by the same audience and writable by nobody. Residual Medium while D-066 and D-070 are open. OQ-197 |  | Residual Medium · OQ-197 |
| 12.2 |  | R-101 Recording validation evidence: A re-test overwrites the earlier result, keeps the first date, or names somebody who did not execute it | One row per execution; executor from the session and shown; the execution time of that execution; a refused save stated; controls offered only to those the database admits. Residual High until D-068 is closed. OQ-200 |  | Residual High · OQ-200 |
| 12.3 |  | R-103 Shared devices and the administrator preview: Records remain on a device after sign-out, or a preview writes under the administrator’s identity | Sign-out removes the offline registers, every screen cache and the menu counts; a preview refuses every write and is recorded. Residual Medium while D-069 and D-070 are open. OQ-204, OQ-205 |  | Residual Medium · OQ-204, OQ-205 |
| 12.4 |  | R-43 Handling of the database credential by automation: A fragment of the credential is written to a log the organisation does not control | The runner refuses before invoking the client when the URL is malformed in the way that causes a fragment to be reported, and prints none of the value; every tail of the password is additionally masked; the credential is passed by environment and never on a command line |  | Residual Low |
| 12.5 |  | R-10 Content integrity: Malicious HTML stored in a KB article | Allowlist sanitiser on save and render; OQ FRS-025 |  | Residual Low |
| 12.6 |  | R-12 Change control: An untested change reaches production | Version control, build ID, changelog, regression tests; periodic review |  | Residual Low |
| 12.7 |  | R-23 Identifier continuity: The identifier counter is left beneath identifiers a migration loaded, and re-issues them | An explicit identifier advances the counter past itself on every write, so a load cannot strand it; a repair function; identifiers already issued out of sequence are re-lettered, not renumbered; OQ-32 |  | Residual Low |
| 12.8 |  | R-24 Bulk load: A load writes NULL into a column the file did not carry, or writes an empty value over a meaningful default | Rows are grouped by column set so an absent column defaults; the constraint fails the batch rather than admitting the row; the loader’s checks reject filling gaps with empty values; OQ-33 |  | Residual Low |
| 12.9 |  | R-25 Response time: A register does not return within a working time on production volume | JIT disabled after measurement; paging; server-side search; PQ-05 measures on full volume with access rules in force |  | Residual Low |
| 12.10 |  | R-31 Quality check before return: Equipment leaves with a failed check, with no check, or with a check nobody was authorised to sign | The check is a separate permission enforced in the database, so it can be held by somebody other than whoever did the work; triggers refuse Ready/Dispatched/Closed on a failed check and refuse a repair or rework dispatched with none. Segregation of the two people is recorded and shown, not enforced, the procedure being silent. OQ-45 |  | Residual Low · OQ-45 |

# Validation Tasks

Each check of 3, with the requirements and tests that prove it. Automated tests run on every change; the last run (27-Oct-2026) passed 200 of 200 database test suites and 22 of 22 automated checks [5]. A manual test is executed and signed here.

|  | Module | Check plan | Responsibility | Validated by | Verification Status |
| --- | --- | --- | --- | --- | --- |
| 1 | Order processing – Spare request | Ability to place a request for spares from engineers, against a call or for hand stock | Service engineer · Reporting Manager · Commercial · NSM |  | Requirements: URS-007, URS-008, URS-021, URS-022, URS-028, URS-035, URS-094, URS-118 … (+16 more) Functional: 53 (FRS-010, FRS-011, FRS-012, FRS-026, FRS-027 … (+48 more)) Automated tests: 39 of 39 passed 27-Oct-2026 Manual tests to execute: OQ-07, OQ-08, OQ-18, OQ-20, OQ-23, OQ-105 … (+22 more) |
|  |  | Line-by-line approval by the Reporting Manager, Commercial and NSM as the cover requires |  |  |  |
|  |  | Re-assignment of a request before it is issued, with a log |  |  |  |
| 2 | Stock maintenance for service spare parts | Stock out by stores to the corresponding engineers, with the Delivery Challan | Service Stores · Spare Coordinator · Service engineer |  | Requirements: URS-008, URS-009, URS-010, URS-021, URS-022, URS-023, URS-024, URS-036 … (+45 more) Functional: 151 (FRS-012, FRS-013, FRS-014, FRS-015, FRS-026 … (+146 more)) Automated tests: 73 of 73 passed 27-Oct-2026 Manual tests to execute: OQ-08, OQ-16, OQ-17, OQ-18, OQ-20, OQ-23 … (+93 more) |
|  |  | Acknowledged receipt per delivery |  |  |  |
|  |  | Hand stock derived from issues, consumption, transfers and returns |  |  |  |
|  |  | Consumption booked against a visit; reconciliation and voiding |  |  |  |
| 3 | Sales – Warranty and Non-warranty | Product sale entry with its warranty | Service admin · Commercial |  | Requirements: URS-010, URS-011, URS-031, URS-040, URS-053, URS-060, URS-068, URS-069 … (+31 more) Functional: 121 (FRS-015, FRS-016, FRS-037, FRS-042, FRS-046 … (+116 more)) Automated tests: 63 of 63 passed 27-Oct-2026 Manual tests to execute: OQ-25, OQ-26, OQ-27, OQ-29, OQ-33, OQ-40 … (+71 more) |
|  |  | The Product Database names the party of the latest sale entry or ownership transfer |  |  |  |
|  |  | Ownership transfer of a machine |  |  |  |
| 4 | Annual Maintenance Contracts | Contract entry, renewal and PM visit count | Service admin · Commercial |  | Requirements: URS-011, URS-048, URS-068, URS-072, URS-076, URS-137, URS-140, URS-141 … (+6 more) Functional: 49 (FRS-016, FRS-056, FRS-079, FRS-080, FRS-084 … (+44 more)) Automated tests: 30 of 30 passed 27-Oct-2026 Manual tests to execute: OQ-27, OQ-42, OQ-64, OQ-106, OQ-107, OQ-118 … (+21 more) |
| 5 | Complaint handling | Breakdown call registration (from a call request, or directly) | Hotline · Customer Support Manager · Service engineer |  | Requirements: URS-002, URS-003, URS-004, URS-014, URS-025, URS-032, URS-034, URS-036 … (+58 more) Functional: 144 (FRS-003, FRS-004, FRS-005, FRS-006, FRS-007 … (+139 more)) Automated tests: 75 of 75 passed 27-Oct-2026 Manual tests to execute: IQ-02, OQ-02, OQ-03, OQ-04, OQ-05, OQ-11 … (+93 more) |
|  |  | Daily complaint review stage 1 |  |  |  |
|  |  | Call allocation |  |  |  |
|  |  | Daily complaint review stage 2 |  |  |  |
|  |  | Visit work details entry |  |  |  |
|  |  | Service report entry with spares used |  |  |  |
|  |  | Service manual report upload |  |  |  |
|  |  | Daily complaint review stage 3 |  |  |  |
|  |  | Field Failure Report raised from a review |  |  |  |
| 6 | Installation & Preventive maintenance | Installation call generation | Service admin · Hotline · Service engineer |  | Requirements: URS-003, URS-005, URS-006, URS-026, URS-040, URS-070, URS-082, URS-093 … (+17 more) Functional: 63 (FRS-005, FRS-006, FRS-008, FRS-009, FRS-030 … (+58 more)) Automated tests: 50 of 50 passed 27-Oct-2026 Manual tests to execute: OQ-03, OQ-04, OQ-06, OQ-14, OQ-19, OQ-21 … (+35 more) |
|  |  | PM call generation (bulk and PM Due) |  |  |  |
|  |  | Installation & PM call visit details and closure report |  |  |  |
| 7 | Customer feedback | Collection of feedback | Service engineer |  | Requirements: URS-012, URS-013, URS-071, URS-109, URS-111, URS-113, URS-114, URS-140 … (+4 more) Functional: 32 (FRS-017, FRS-018, FRS-074, FRS-083, FRS-108 … (+27 more)) Automated tests: 26 of 26 passed 27-Oct-2026 Manual tests to execute: OQ-10, OQ-28, OQ-54, OQ-96, OQ-100, OQ-123 … (+16 more) |
| 8 | Indoor service (workshop) | Intake, cleaning, work, quality check and release of equipment in the workshop | Service engineer (workshop) · QC |  | Requirements: URS-050, URS-051, URS-116, URS-117, URS-169, URS-170, URS-171, URS-173 … (+3 more) Functional: 19 (FRS-057, FRS-058, FRS-059, FRS-143, FRS-144 … (+14 more)) Automated tests: 11 of 11 passed 27-Oct-2026 Manual tests to execute: OQ-43, OQ-44, OQ-45, OQ-136, OQ-206, OQ-222 … (+3 more) |
|  |  | Pre-Delivery Testing and recycled spares |  |  |  |
| 9 | Masters, users and access | Party, user and master-list maintenance | Administrator · Service admin |  | Requirements: URS-002, URS-010, URS-014, URS-056, URS-063, URS-067, URS-074, URS-078 … (+34 more) Functional: 146 (FRS-001, FRS-002, FRS-003, FRS-004, FRS-014 … (+141 more)) Automated tests: 70 of 70 passed 27-Oct-2026 Manual tests to execute: IQ-02, OQ-01, OQ-02, OQ-09, OQ-11, OQ-22 … (+94 more) |
|  |  | Roles & permissions; a login only for a person on the User Master |  |  |  |
|  |  | Signatures, held for and placed by their owner |  |  |  |
| 10 | Training and controlled documents | Training assignment, sessions and Training Form R/HR/11 | VP Technical · Engineer R&D and QMS |  | Requirements: URS-020, URS-029, URS-030, URS-057, URS-079, URS-099, URS-102, URS-125 … (+11 more) Functional: 40 (FRS-025, FRS-035, FRS-036, FRS-066, FRS-067 … (+35 more)) Automated tests: 15 of 15 passed 27-Oct-2026 Manual tests to execute: OQ-13, OQ-21, OQ-51, OQ-106, OQ-110, OQ-114 … (+21 more) |
|  |  | QMS documents and service manuals |  |  |  |
| 11 | Reports, KPI and quality objectives | KPI, objectives and dashboards computed by the database | NSM · Customer Support Manager |  | Requirements: URS-013, URS-014, URS-019, URS-036, URS-039, URS-040, URS-045, URS-054 … (+31 more) Functional: 111 (FRS-009, FRS-018, FRS-019, FRS-024, FRS-035 … (+106 more)) Automated tests: 56 of 56 passed 27-Oct-2026 Manual tests to execute: IQ-01, OQ-10, OQ-11, OQ-14, OQ-21, OQ-29 … (+82 more) |
|  |  | Reports and exports |  |  |  |
|  |  | Bulk uploads of historical data |  |  |  |
| 12 | Data storage, audit and security | Audit log and record retention; quality records never deleted | Administrator · QA |  | Requirements: URS-016, URS-019, URS-084, URS-090, URS-128, URS-129, URS-149, URS-157 … (+5 more) Functional: 67 (FRS-013, FRS-021, FRS-024, FRS-070, FRS-091 … (+62 more)) Automated tests: 37 of 37 passed 27-Oct-2026 Manual tests to execute: IQ-01, OQ-08, OQ-09, OQ-53, OQ-90, OQ-91 … (+44 more) |
|  |  | Row-level security; nobody signed out or unknown holds anything (security scan) |  |  |  |
|  |  | Backups and validation records |  |  |  |

Non-auditable requirements (12) are listed in the package [4] and are not offered as evidence against a clause. Requirements in the package: 189 user, 285 functional; tests: 313. Open defects in the register: 87 (D-001, D-002, D-003, D-004, D-005, D-006, D-007, D-008, D-009, D-010 … (+77 more)).

# CONCLUSION

RITHI is validated for its suitability and all mitigation measures have been implemented and followed, subject to the open items recorded in the package [4] and in the Security Scan report [6].

Verified by the responsible persons

Validation Date:                                        Validated by:

END OF DOCUMENT